## Slide 1: High-Z Sensor Development at DECTRIS
Pietro ZAMBON
20.08.2023
dectris.com

## Slide 2: Table of contents
CdTe
GaAs:Cr
CZT
Perovskite
04
02
03
01
All results obained in COUNTING MODE

## Slide 3: CdTe
Commercially available, reliable supplier (Acrorad, JP)
Large sensor area up to 4 cm x 4 cm
Ohmic contacts, 750 um and 1000 um sensor thickness
75 um, 100 um and 150 um pixel size
Electron collection
Ileak 2-3 uA/cm2
Settling time ~15-20 min at bias ON
Generalities

## Slide 4
PILATUS3 CdTe 2M @ DanMAX beamline, MAX IV
Beam energy 15-35 keV
Powder X-ray diffraction

## Slide 5
Spectral properties
Refs:
Zambon, et al., Spectral response characterization of CdTe sensors of different pixel size, NIM A 892 (2018).
Zambon, et al., A fitting model of the pixel response to monochromatic X-rays in photon counting detectors, NIM A 905 (2018).
Element | Kedge | Kα1 | Kα2 | Kβ1
Cd | 26.73 | 23.17 | 22.99 | 26.09
Te | 31.82 | 27.47 | 27.20 | 30.99
X-ray 40 keV

## Slide 6
Spectrum = derivative of Vcmp-scan = ramp function convoluted with a Gaussian (noise).
Direct X-ray tube beam given by Bremsstrahlung radiation. It decreases linearly to the cut-off point.
For the Duane-Hunt law it corresponds to the tube kVp.
spectrum’
2.
spectrum’’
3.
4.
Spectrum’ = step function convoluted with a Gaussian
The peak is the calibration point
Spectrum’’ = delta function convoluted with a Gaussian = Gaussian.
Ideal
Measured
spectrum
1.
end-point
Ref:
Zambon, et al., A wide energy range calibration algorithm for X-ray photon counting pixel detectors using high-Z sensor material, NIM A 925 (2019).
High-energy calibration

## Slide 7
Beyond X-ray fluorescence: 99mTc spectrum
ACCURACY: ± 0.5 keV (< 1%) → global threshold calibration works!
PRECISION: Pixel-to-pixel threshold dispersion ~1.7 keV
Flat-field Image at Eth=120 keV
Tc source
Pb shielding
Setup
1 mm CdTe 150 µm pixel size Calibration range: 25 – 140 keV
1 cm
Notes: γ = 140.5 keV (98%) + 142.6 keV (1.4%)

## Slide 8
Pilatus3 flat-field (Ag, 50% threshold)
Network of lines i.e. local charge build-up with distorting effects on charge transport.
dark+bright line, constant average counts
depends on growing process
enhanced by time (h) and irradiation
mitigated by high Vbias
erased (almost completely) by detector reset
Ling 2013
Very small number of defects and/or bad pixels
Excellent dicing edge quality

## Slide 9
Measurement of the count rate vs. bias voltage, tube voltage, tube current, time (and temperature)
Polarization example in CdTe
-200V
-400V
-100V
-300V
-500V
direct beam 60 kVp, tube current increased every 15 min
bias -500 V, tube current increased every 2 min
Count rate degradation with increasing time, flux, energy and for decreasing temperature (not shown). Uniform degradation over the irradiated area.

## Slide 10: GaAs:Cr
Generalities
Cr-compensation process firstly developed at Tomsk University, now available also at Dectris.
Reached a good manufacturing level, still room for optimizations
Polarization and grains seem to be related with GaAs starting material (doping, stresses...) and the Cr diffusion
6’’ wafer LEC, 4’’ wafer VGF
Large sensor area up to 8 cm x 4 cm
Ohmic contacts, sensor thickness from ~250 um to 1 mm
Electron collection
Ileak min. ~5 uA/cm2
No settling time at bias ON

## Slide 11
Sensor edges (2 pixels size)
Grainy structure lead up to a factor 3 difference in count rate. Strongly dependent on the starting material
Example of flat field image
Full sensor 8 cm x 4 cm, 500 um-thick 75 um pixel size
1 cm
LEC
VGF

## Slide 12
X-ray tube 28 kVp with Al filter, Eth = 8 keV
LEC material, Flat-field corrected, gaps interpolated.
1 cm

## Slide 13
GaAs shows less fluorescence yield
Comparable energy resolution
Spectra comparison GaAs/CdTe
X-ray 60 keV
X-ray 40 keV
normalized to the QE
normalized to the QE
E0
CdKα
TeKα
E0-CdKα
E0-TeKα
E0
CdKα
TeKα
E0-CdKα
E0-TeKα
GaKα
AsKα
E0-GaKα
E0-AsKα
GaKα
AsKα
E0-GaKα
E0-AsKα
Element | Kedge | Kα1 | range (μm)
Ga | 10.39 | 9.25 | ~40
As | 11.88 | 10.54 | ~15

## Slide 14
GaAs sensor: 75 μm pixel, 500 μm thickness, room temperature
Count rate capabilities
X-rays 55 keV
X-rays up to 20 keV
Strong polarization for increasing energies!

## Slide 15
Direct X-ray tube beam 60 kVp, threshold 17 keV
Example of polarization behavior
Increasing tube current
Low rate
Good
Bad
MAX observed count rate: 0.6-1.5 Mcts/s/pix
i.e. 105-266 Mcts/s/mm2
1 cm

## Slide 16
Tomsk GaAs (diode), dedicated setup at Dectris. Mo X-ray fluorescence spectra (17.48 keV), range ~30 um
Charge collection efficiency (CCE) measurements
Hecht equation only electrons
Hecht equation electrons + holes:
µτe = 3.74·10-4 cm2V
µτe = 3.83·10-4 cm2V
µτh = 1.69·10-6 cm2V

## Slide 17
Assuming
µe = 1500 cm2/V/s
µh = 100 cm2/V/s
CCE results summary
sample | µτe (cm2 V) | τe (ns) | µτh (cm2 V) | τh (ns)
ArN276_G555715_1_N13 | 4.21 · 10-4 | 280 | 8.72 · 10-6 | 87
ArN204_614378_12_N2 | 1.51 · 10-4 | 100 | 5.29 · 10-6 | 53
Ag4669A | 2.37 · 10-4 | 158 | 4.06 · 10-6 | 40
Ag4668C | 3.02 · 10-4 | 201 | 4.86 · 10-6 | 48
Ag4669D | 2.26 · 10-4 | 150 | 3.61 · 10-6 | 36
Holes lifetimes consistent with Ilahi, Doping effects on minority carrier parameters in bulk GaAs, Phys. B 652 (2023)
And used in Zambon, Simulation of polarization dynamics in semi-insulating, Cr-compensated GaAs pixelated sensors under high x-ray fluxes, AIP Adv. 11 (2021)

## Slide 18
Direct X-ray tube 23 kVp
Eth = 10 keV
Pixel size 75 um
Geometrical distortion due to the grainy structure
Pinhole scan
5 um step

## Slide 19
Goal:
On the basis of a flat-field illumination only, infer the geometrical distortion assuming pixels as irregular quadrilaterals.
The area of the quadrilaterals corresponds to the number of counts in that pixel → compute optimal coordinates of the vertices (non-linear, huge).
Is a correction possible?
Reconstructed boundaries
Ideal boundaries
Reconstructed centroids
Ideal centroids
Measured centroids

## Slide 20: CZT
Generalities
Commercially available (Redlen)?
Typically sensor area 2 cm x 2 cm
Ohmic contacts, 1-2 mm sensor thickness
75 um, 100 um and 150 um pixel size
Electron collection
Ileak 1-2 uA/cm2
No settling time at bias ON
“High Flux” type → very resitent towards polarization

## Slide 21
Reached very good response uniformity
Sometimes defects can be found (bulk, surface)
Dicing edges not always excellent: no visible damages but bad pixel response
Examples of flat field images (X-ray fluorescence)
1 cm

## Slide 22
Count rate w/ KITE ASIC, non-paralyzable (retrigger enabled)
7 keV
17 keV
43 keV
Electrons 300 keV – Retrigger ON
Threshold (keV) | Cutoff at slope 0.2
7 | 60.6 Mcts/s/pix
17.5 | 84.0 Mcts/s/pix
43 | 126.1 Mcts/s/pix
Ref:
Zambon, et al., KITE: High frame rate, high count rate pixelated electron counting ASIC for 4D STEM applications featuring high-Z sensor, NIM A 1048 (2023)

## Slide 23
Bipolar signals are induced in neighboring pixels (Ramo’s theorem)
The fast shaping time of the front-end electronics (< 10 ns FWHM) makes them measurable
Spectra (electrons) and bipolar signals w/ KITE ASIC
60 keV
120 keV
140 keV
160 keV
200 keV
100 keV
80 keV
Negative peaks

## Slide 24: PEROVSKITE
Sample sensors from Actinia https://actiniadetectors.com/
Inorganic Pb-based (CLB) Perovskite: CdPbBr3 development since 2000s, bandgap ~2.3 eV (-> low leakage current)
Sample area ~1 cm x ~1 cm, 150 um pixel size
Generalities
PRELIMINARY

## Slide 25
First images with Mo fluorescence (~17.5 keV) with both high voltage polarities
First batch Dec 2022
Second batch Aug 2023
Vbias +600 V
HOLE collection mode
Vbias -600 V ELECTRON collection mode
X-ray illumination
1 cm
0 counts

## Slide 26
Hole collection mode
Mo Fluorescence spectra
Electron collection mode

## Slide 27: Conclusion
High-Z sensor development is an open research topic at DECTRIS since 10+ years
Characterizing/ed CdTe, GaAs, CZT, Perovskite in photon/electron counting mode
Capable of manufacturing GaAs:Cr sensors
A universal high-Z sensor does not exists. The application matters!
Aknowledment
V. Radicci, S. Fernandez Perez, G. Montemurro, M. Rissi, A. Jensen, A. Gonzalez Taboada, T. Sakhelashvili.

## Slide 28: Thank you for your attention!
DECTRIS Ltd.
5405 Baden-Dättwil
Switzerland
www.dectris.com