## Slide 1: News from the XOG: recent achievements & challenges
Ray Barrett (XOG leader)
B. Argoud, P. Bras, J. Cave, J-P. Guigay, S. Laboure, C. Lapras, T. Manning, C. Morawe, F. Perrin, B. Picut, T. Roth, T. Sieg-Letessier, R. Verbeni, A.Vivo
Multilayers
Refractive lenses
Crystals
Crystal analysers
Metrology
‘Routine’ manufacture/test
Developments
  Wavefront preservation
  Optics reconditioning
  Optical systems new/improved capabilities

## Slide 2: Multilayer Lab (ML): Introduction
- 2 magnetron sputter deposition systems
Graded MLs for monochromators and KB focusing mirrors
Dynamic deposition technique
Substrate
Sputter Target
Mask
Large ML Coating System – LMCS (2007): up to 1 m long optics
Compact ML Coating System - CMCS (2021): up to 0.3 m long optics
Particle flux

## Slide 3: XOG ML Lab: Main achievements
- Nano KB coating for ID16A using CMCS
- Smoother carrier motion during deposition
- Improved graded d-spacing control compared to LMCS: dΛ/Λ  0.1%

## Slide 4: XOG ML Lab: Challenges
2D static profile
INPUT
COMPUTATION
OUTPUT
Targeted thickness profile (2D)
Static flux profile (2D)
Deep learning (neural network)
Slit shape
Speed profile
Slit
Speed profile
- Advanced mirror surface figure correction using differential deposition (DD)
- DD with in-situ metrology tools for faster convergence
Shack-Hartmann
wave front sensor
differential deposition
- DD in 2 dimensions using deep learning

## Slide 5: CRLs: Main Achievements
start of 1D Al lens manufacture:
First lens pressed in January ’25
Total pressed: 51 so far
manufacture of 2D Al lenses:
N=173 R=30um to ID11
N=70 R=30um to ID01
N=24 R=30um to ID06-LVP
ID06-EH1 operational
1D lens press
operating on BM05
Aluminium
Beryllium
significantly reduced tails
ISDD lenses operating on ID01 @ 20keV
280nm beam
FWHM

## Slide 6: CRLs: Challenges
Delivery of R=50µm 1D Al lenses to beamlines
shape fidelity, punch aging
height error [µm]
Delivery of R=30µm & R=200µm 2D Al lenses to beamlines
Phase correction optics
R. Celestre re-joining XOG for these efforts

## Slide 7: Crystal Lab: Achievements
2025 production
12 x monochromator crystals for the new DCMs of ID16B/ID26 beamlines
2 x backscattering monochromators for ID28 beamline
4 x monochromator crystals for CRG’s and external companies
2025 developments
New X-ray Laue Camera for crystal orientation
DDCOM commissioning: precise asymmetry angles ~0.001 degrees
Second LamPlan polishing machine
ID28 Backscattering
Bonse-Hart LLL-type interferometer crystal
Laue picture Si(111)
40 s exposure
New X-ray Laue camera
DCM monochromator crystals
DDCOM machine
Precise asymmetry angle measurements
Second LAMPLAN
Polishing machine

## Slide 8: Crystal Lab: Challenges
Future perspectives:
Improving the planarity and roughness of the polished crystals
Implementation of pitch* polishing
* pitch = mixture of rosin and bitumen
(rosin = distillate of wood resin)
B. Argoud and B. Picut during “pitch polishing” training in Paris
Pitch polishing plate
New inhouse designed X-Y polishing machine dedicated
to pitch polishing
pieplow-brandt.de

## Slide 9: Crystal Analyser Lab: Achievements
2025 Production:
25 x bent, bent-striped and diced spherical analysers
16 x cylindrical Von Hámos analysers
5 x vacuum clamped cells (spherical/cylindrical)
~50% for external clients
Improvements/developments
  Cylindrical VCC
  Spin-coating technique
Bent
Bent-striped
Von Hámos
Diced
VCC Spherical
VCC Cylindrical
Typical ESRF 5-crystal spectrometer

## Slide 10: Crystal Analyser Lab: Challenges
Challenges for the next years
Find alternative Ge ingot suppliers
Improvement of bonding of Ge analyser crystals
  epoxy gluing
Production of quartz and lithium niobate spherical/cylindrical analyser crystals
Find new suppliers of double-side polished wafers
Ge ingots very important for analyser manufacture
Major suppliers in China (control 80-85% of world production). Export controls implemented in 2023-24. Export bans to US - very difficult to purchase.
Price increased by 4X in one year.
Spin-coating machine:
Epoxy gluing of spherical/cylindrical analyser for Ge/quartz and LiNbO3
Improvement of the spinning parameters.
More adapted Epoxy glue
Dicing LiNbO3 appears to be under control.
Dicing quartz wafer still very hard  few analysers in the world
Alpyx very slow delivery
Very difficult to find suppliers for ‘non-standard’ reflections
Double-side polished Si wafers
Siltronix
Alpyx
Diced LiNbO3 wafer – Ø65 mm

## Slide 11
Metrology Lab: Achievements
Measurement of Sub-nanometer deformation induced by support
Mirror support allowing various sphere position
  Optical surface
  Optical surface
  Top view
  244 mm x 20 mm
  0.7
  -0.7
  nm
  Optical surface Shape error
JTEC mirror (PSI): 270 mm x 40 mm(w) x 50 mm (t)
Face Up/Face down measurement capability successfully implemented on Fizeau interferometer
Fizeau
Folding mirror
Mirror facing up
LEAPS-Innov Superflat prototype mirror: 600 mm x 60mm x 60 mm
Face side
Face side
Face Up
  8
  -8
  nm
500 mm x 20mm
LTP scanning speed successfully increased to 100 mm/sec

## Slide 12
Metrology Lab: Challenges
Metrology
Improve gluing process for ID16B nano-KB
Improve mirror support for acceptance tests
Reduce low spatial frequency measurement errors (LTP/ Fizeau) –air turbulence
Improve micro-stitching interferometry for longer mirrors (currently <150 mm)
    new micro-interferometer arriving in 2026
Mirror refurbishment
strip off coatings and recover original Si surface
remove contamination after X-ray exposure
metrology for 2D differential deposition aimed to improve mirror figure errors
Mirror implementation
Reduce or compensate mirror mounting deformations (FEA as for gravity : ID32 – ID29 …) by polishing or differential deposition ??
ID16B
Nano KB
Before bonding
ID21 HF
ID21 VF
HF
VF
After O2 plasma exposure

## Slide 13
Thank you!