## Slide 1: APPLE III developments at DESY
Variable polarisation for FLASH
Kathrin Götze (DESY Undulator group)
ID26 workshop Deauville, 15th May 2026
On behalf of: Markus Tischer, Pavel Vagin, Andreas Schöps, Paul Neumann, Thorsten Vielitz, Philip Eckoldt and the other members of the DESY undulator group, and with thanks to colleagues at the central mechanics and mechanical design groups and at FLASH machine and controls

## Slide 2: Outline
APPLE III developments at DESY
APPLE III concept
The “prototype”: APPLE III Afterburner for FLASH2, UE17.5
The seeding upgrade: Six UE35 APPLE III for FLASH1
Lessons learnt
Summary and outlook
e-

## Slide 3: APPLE III concept
  APPLE III: highest field, less bulky, cost efficient
  Force reduction up to factor of 8 - reuse of planar undulator support structures
Gap change
e-
A
B
AB
Shift drive for different polarization modes
shift λ/2
shift λ/4
shift 0
Variable field direction:
APPLE III design: Bahrdt J, Frentrup W, Gaupp A, Kuske B, Meseck A, and Scheer M 2004 Proc. SRI Conference 2003 AIP CP 705 215‒18
Source shift drive picture: Hannon, F.E., et al., Proceedings of EPAC. 2004.
Variable polarisation for investigating magnetism and chirality
Notes: Force compensation allows to reuse our usual planar undulator support structures (e.g. in comparison to bulky PU04 without compensation, or APPLE-X)

## Slide 4: APPLE III concept
Magnet and keeper design
◄FLASH2 Afterburner keeper 17.5mm period length
Field tuning by virtual shimming
Keeper body: aluminium, wedge: copper tin bronze CuSn2.
Differential screw (M3/M4 thread)
One full period per keeper, separated by cut in keeper body for individual tuning of half periods.
Function magnets
Force compensation magnets
Force compensation magnets
Differential screw
Wedge
Pretension Screw and Spring
Vacuum chamber:
FL1 radiator keeper ► 35mm period length
8
mm
blade
0.5mm
1mm
magnetic gap
Full period (17.5 or 35mm) keepers

## Slide 5: APPLE III concept
Magnet and keeper characterisation
Sorting list for each device based on these data
Assembly onto girders/support structures
Measurement and tuning of complete device with Hall probe and stretched-wire set-up
Stretched-wire and Hall measurements of all individual keepers
Glued AB magnets measured individually by Helmholtz coil
Magnets sorted into pairs and mounted into keepers
(one keeper=one period=two function magnet pairs)
◄▲Keeper on measurement stand
Measurement stand scheme

## Slide 6: APPLE III Afterburner
Circularly polarised SASE radiation at FLASH2.
Period length 17.5mm
Wavelength range FLASH1: 3.3 – 90 nm (3rd Harmonic Wavelength down to 1.7 nm)
Radiates on the third harmonic of the SASE section to provide photon energies around the L-edges of iron, cobalt and nickel (890-700 eV or wavelength range 1.33--1.77 nm) for studies of ultrafast magnetisation dynamics.
Installed August 2023
First publication May 2025
https://doi.org/10.1063/4.0000298
Notes: Support structures look the same – thanks to force compensation concept, no reinforcement needed

## Slide 7: APPLE III Afterburner for FLASH2
Characteristics and results
<6deg after tuning.
Taper-corrected: <5deg.
K value
Phase error
minimum gap: 1.5mm
see also Markus Tischer et al. 2025 J. Phys.: Conf. Ser. 3010 012045
DOI 10.1088/1742-6596/3010/1/012045
Trajectories (@1.35GeV)
µm
µm
m
m
m
RMS phase error gap dependence
vert
hor
circ
circ
hor pol
ver pol
circ
vert
hor
circ

## Slide 8: APPLE III Afterburner for FLASH2
Characteristics and results
Dipoles: ±50 Gcm gap dependence plus shift-dependent kick error of similar size.
Corrected in a feed-forward by small air coils.
Only small quadrupole contribution.
see also Markus Tischer et al. 2025 J. Phys.: Conf. Ser. 3010 012045
DOI 10.1088/1742-6596/3010/1/012045
Kick errors and multipoles
Transverse dependence of vert. and horiz. field integrals at gap = 2mm
Shift dependence
Gap dependence

## Slide 9: FLASH1 Seeding upgrade
General layout of the radiation section
2 Modulator undulators with 84mm period length and hybrid (magnets&poles) structure.
3 refurbished planar U32 undulators from sFLASH to increase the pulse energy at short wavelengths.
6 radiator undulators of APPLE III type UE35.
Wavelength range: 4 – 60 nm
Radiator undulators (APPLE III)
U32 undulators from sFLASH
e-beam

## Slide 10: Six UE35 APPLE III for FLASH1
Characteristics and results
K value
minimum gap: 1.3-1.4mm
Phase error
Kmax=4.3
Significant variation of phase error for different polarisation modes. Correction by taper and resistive long coil.
without taper& long coil
with taper& long coil
Trajectories (@1.35GeV)
vert
without taper correction & long coil
with taper correction & long coil
hor
circ
vert
hor
circ
ph.err, deg
ph.err, deg
x, mm
x, mm
x, mm
RMS phase error gap dependence
Gap dependence
vert
hor
circ
vert
hor
circ
Notes: coil is 3 turns (at bottom and top, respectively) of ~5mm^2 copper wire
makes 38uT per Ampere at minimum gap 1mm, and 6.3uT at fully opened gap 200mm
x2.5m long makes 95mTmm/A @ 1mm and 16mTmm/A @ 200mm
maximum current 10Amp -> 380uT / 0.95Tmm
coil width 100mm, distance between top&bottom coils 50mm at minimum gap

## Slide 11: Six UE35 APPLE III for FLASH1
Characteristics and results
Multipoles errors
Transverse dependence of vert. and horiz. field integrals for shift 0 (hor pol) and shift 0.5 (circ) at different gaps
Optimized for circular mode
Optimized for circular mode
Shift and gap dependence of vert. and horiz. quads
I1, Tmm
I1, Tmm
hor
circ
hor
circ
I1V
I1H
I1V
I1H
ver dipole, shift dep
hor quad, shift dep
ver quad, shift dep
hor dipole, shift dep
vert
hor
circ
vert
circ
vert
hor
circ
vert
circ
shifts
Notes: coil is 3 turns (at bottom and top, respectively) of ~5mm^2 copper wire
makes 38uT per Ampere at minimum gap 1mm, and 6.3uT at fully opened gap 200mm
x2.5m long makes 95mTmm/A @ 1mm and 16mTmm/A @ 200mm
maximum current 10Amp -> 380uT / 0.95Tmm
coil width 100mm, distance between top&bottom coils 50mm at minimum gap

## Slide 12: Lessons learnt
Gluing of magnets to AB-pairs
One-period keepers for UE35 (and for UE17); longer super-structures (e.g. like SwissFel Apple-X) could be good option for future. Single keeper flipping/swapping was not necessary.
Pre-characterisation at several stages in the assembly process (magnets separately, single keepers, sub girders mechanically) costs time but is worth it.
Additional linear encoder for shift drive regarded as mandatory; helped to detect deficiencies with the shift drive.
Shift drive design needs substantial safety margins in terms of load. Exceeding manufacturing tolerances will quickly use up any buffer on the parameters of the drive train assembly (motor, coupling, bearings, spindle nut).
Focusing properties of Apple-3/X structures: need sufficient awareness on the beam physics side.
AB

## Slide 13: Lessons learnt
“Planar” Hall effect: Hall measurements show an unusual behaviour of the transverse field for the U84 modulators
  Large discrepancy between Hall probe and stretched-wire measurements for transverse direction that cannot be explained by misalignment, mispositioning or wire sag. Adjusted the Hall calibration polynomial by Bz+=0.00011By2 (dashed). But: correction should be checked/recalibrated by stretched wire for different devices.
  If sensor is rotated by 90 deg around horizontal axis, difference to the wire is significantly reduced. → Horizontal sensors are vertically displaced and there is no vertical gradient of the vertical field.
https://www.senis.swiss/wp-content/uploads/2024/10/Datasheet-F3A-Magnetic-Field-Transducer_rev.1.2.pdf
1.5
0.5
1.0
0.0
10
20
50
100
200
Field(T)/Field integral(Tmm)
Undulator gap (mm)
0.35
Tmm
beam direction
Notes:   "with adjusted calibration polynomial by Bz+=0.00011By^2" correction should be checked/recalibrated by stretched wire for different devices (most probably depends on undulator period/field gradients)
  If sensor is rotated by 90 deg around horizontal axis, difference to the wire is significantly reduced.
  Because in this case the horizontal sensors are vertically displaced and there is no vertical gradient of the vertical field.

## Slide 14: Lessons learnt
shift moving from
-1…….-0.5..…..0..…..+0.5.……+1
Skew quad investigation:
Conducted optical measurements of local mechanical movement of the magnets during shift movement (resolution <1um)
Magnets on one keeper move by about 25um and 0.1mrad in pitch angle during one shift cycle.
Difference between magnet pairs on one keeper: 5-7um and 1E-5rad.
Small but systematic difference for "positive" / "negative" halfperiods most probably causes multipoles shift dependence.

## Slide 15: Lessons learnt
Skew quad investigation:
Conducted optical measurements of local mechanical movement of the magnets during shift movement (resolution <1um)
Magnets on one keeper move by about 25um and 0.1mrad in pitch angle during one shift cycle.
Difference between magnet pairs on one keeper: 5-7um and 1E-5rad.
Small but systematic difference for "positive" / "negative" halfperiods most probably causes multipoles shift dependence.
Shift, Pi
Shift, Pi

## Slide 16: Summary and Outlook
We successfully designed, manufactured and tuned six APPLE III UE35 and one afterburner APPLE III UE17 for the free-electron laser FLASH.
Devices were tuned by virtual shimming (wedge movement on the keepers), magic fingers, (shift dependent) taper and resistive long coils.
Phase error between 2 and 10 deg for UE35 and <5deg for UE17 afterburner.
Hall measurements show an unusual behaviour of the transverse field for very high vertical fields (gradients), calibration function needs to be adapted.
UE35: Skew quads minimized for circular mode. Might be caused by keeper deformation/movement of halfperiods. Further investigation by optical measurement of magnet displacement during shift movement planned.
Notes: UE17 no skew quad

## Slide 17: Back up
Additional slides

## Slide 18: The FLASH2020+ upgrade
The future high repetition rate XUV and soft X-ray FEL facility
Small increase in electron beam energy to 1.35 GeV
Tunable undulators at FLASH1 allow fully parallel operation of two FEL lines.
FLASH1 fully externally seeded with full repetition rate that FLASH can provide in burst mode.
Afterburner for FLASH2 to provide circularly polarised light
Extended the wavelength reach of the fundamental harmonics to the oxygen K-edge and L-edges of Fe, Co, and Ni.
FL1: 4 – 60 nm, FL2: 3.3 – 90 nm (3rd Harmonic Wavelength down to 1.7 nm)

## Slide 19: Six UE35 APPLE III for FLASH1
Mechanical shimming
  Manual shimming of subgirders to even out differences in height of up to 200um (“chain of tolerances”)
Support structures
Horizontal profile before (dashed) and after shimming (Touch probe)
0.3
0.0
0.1
0.2
-0.1
-0.2
-0.3
-1000
-500
500
1000
0
x, mm
h, mm
Left
Right
Position of shims for vertical and horizontal shimming
Notes: Touch probe measurements of the girder profiles after assembly to the support structure but before mounting keepers, with dummy mechanical keeper to measure both actual horizontal and vertical positions. Hall probe bench temporarily converted into CMM.

## Slide 20: Six UE35 APPLE III for FLASH1
Characteristics and results
After multipole tuning and magic fingers: ±0.15Tmm (ver), ±0.25Tmm (hor)
Skew quad optimised for circular mode
K value
minimum gap: 1.3-1.4mm
Phase error
Kick errors and multipoles
Kmax=4.3
Significant variation of phase error for different polarisation modes. Correction by taper and resistive long coil.
without taper& long coil
with taper& long coil
before
gap dependence
Notes: coil is 3 turns (at bottom and top, respectively) of ~5mm^2 copper wire
makes 38uT per Ampere at minimum gap 1mm, and 6.3uT at fully opened gap 200mm
x2.5m long makes 95mTmm/A @ 1mm and 16mTmm/A @ 200mm
maximum current 10Amp -> 380uT / 0.95Tmm
coil width 100mm, distance between top&bottom coils 50mm at minimum gap

## Slide 21: Lessons learnt
Skew quad investigation:
Conducted optical measurements of local mechanical movement of the magnets during shift movement (resolution <1um)
Magnets on one keeper move by about 20um and up to 1mrad in angle during one shift cycle.
Difference between magnet pairs on the keeper: 2-3um.
Small but systematic difference for "positive" / "negative" halfperiods most probably causes multipoles shift dependence.
shift moving from
-1…….-0.5..…..0..…..+0.5.……+1
change in position (mm)
change in angle (rad)
difference
magnet pair 1
magnet pair 2
magnet pair 1
magnet pair 2

## Slide 22: APPLEIII Afterburner for FLASH2
In operation
Ball chamber instrument
Measuring angular distribution of photoelectron emission of a certain photoelectron feature from 3rd harmonic
P1= 0.98
P1= 0.99
ABU shift = 0.0
P1= 0.08
P1= 0.98
P1= 0.06
EACC = 880 MeV
hv = 104.6 eV / 313.8 eV
ABU shift = +/- 0.25
ABU shift = +/- 0.5
Circularly polarized light down to 1.46 nm generated
First successful user experiments
Courtesy M. Braune, K. Tiedke, J. Rönsch-Schulenburg and colleagues

## Slide 23: Modulators
Two U84 planar undulators with hybrid structure (magnet and poles)
84mm
Sorting of magnets based on novel Hall mapper reduces amount of pole tuning
Parameters were within specs without any pole tuning (straight trajectory, phase error <4°).
Additional girder curvature spoiled phase error
Strong fields
Strong forces
Reaching limits of sensor calibration
phase
K
5E-5
0
-5E-5
5E-5
0
-5E-5
1.005
1.000
0.995
0.990
10
5
0
-5
-1.5 -1.0 -0.5 0 0.5 1.0 1.5
-1.5 -1.0 -0.5 0 0.5 1.0 1.5
-1000 -500
0 500 1000
x/mm
x/m
Gap/mm
12
8
4
0
Length = 2.5m
Period = 84mm
Min. gap = 9mm
Bmax = 1.8T
Specs: K=11, achieved: K = 12.9

## Slide 24: Modulators
Two U84 planar undulators with hybrid structure (magnet and poles)
K
phase
5E-5
0
-5E-5
5
0
-5
-1.5 -1.0 -0.5 0 0.5 1.0 1.5
After tuning
Trajectories further straightened
0.5deg RMS phase error at minimum gap
Girder shape strikes back at intermediate gaps due to different gap dep.
K = 12.9
„Target" value was Ktarget=11 at 9.5mm gap, to reach 343nm at 1.4GeV
Now at 1.35GeV: 343nm reached at K=10.6, or operating gap ~12.5mm
Both Modulators tuned and close to ready.
-1000 -500 0 500 1000
x/m
x/mm
Gap/mm

## Slide 25: Refurbished U32s from sFLASH
Three devices for the interim to increase pulse energy at short wavelengths
-- Field before
-- Field after flipping 18 magnets US
sfund01
K
phase
Three planar devices with 32mm period length from the former sFLASH experiment were retuned.
sfund01: Severe radiation damage (>10%) – flipped 18 magnet pairs and replaced end magnets upstream.
K between 2.65 and 2.70 at min gap of 9mm
Phase errors <3°
Remaining kick errors: ±50 Gcm. Corrected in a feed-forward by small air coils.

## Slide 26: Resistive coils with variable field direction
Ferrite enforced steerer/corrector coils (air-cooled)
vertical
horizontal
quadrupole
Correct (gap dependent) kick up and down stream of undulators
In-house design for stronger coils with slow feedback option
Resistive, air-cooled coils
Compact and cost-efficient design
Ferrite-based – 0.55Tmm at 1A
Four sub-coils per unit – variable field direction
AC capability
Series of 40 coils manufactured
On-axis first field-integral as a function of current (hor/vert field).
35mm
Stored coils

## Slide 27: Phase shifters
Compact, permanent magnet-based phase shifters on intersections
beam
Magnet blocks. Size: 35x35x15mm3
Phase advance as a function of gap.
All 10 phase shifters achieve at least 2.8*2π for 20nm@1.35GeV.
As measured
Background subtracted
Intersection
Phase shifter
Magnetic set-up
Compact, permanent magnet-based design
Pre-sorting allows for using lower quality, low price magnets
Series of 10 phase shifters built and tuned
<0.02Tmm (6µrad) on-axis kick remaining
high uniformity within series

## Slide 28: Phase shifters and corrector coils
Managing cross-talk
Originally only 22mm distance between permanent magnets of phase shifter and ferrite enforced coils
Strong, unwanted dampening of PS fringe fields due to cross-talk
Increased the distance to 30 mm (max)
Passive compensation by installing two sets of ferromagnetic screws above and below the beam pipe.
Phase shifter
PM
PM
Coil w/ ferrite core
ferromagnetic screws