## Slide 1: Mobile Measurement Systems for Magnetic Characterization of Superconducting Insertion Devices
Andreas Grau, Institute for Beam Physics and Technology (IBPT)
15.05.2026, Insertion Devices Workshop 2026, Deauville, France
Andreas Grau for A. Grau, N. Glamann, B. Krasch, D. Saez de Jauregui,
Karlsruhe Institute of Technology, Karlsruhe, Germany
and
T. Gerhard, A. Hobl, A. Vatagin, W. Walter
Bilfinger Nuclear & Energy Transition GmbH, Würzburg, Germany

## Slide 2: Outline
  Introduction
  Magnetic characterization facilities (MCF) at KIT
  Characterization Setups for Field Error Reduction (CASPER I + II)
  Mobile magnetic measurement setup for field integrals
  Instrumentation and measurement results
  Mobile magnetic measurement setup for local field and field integral characterization of superconducting insertion devices
  Instrumentation and measurement results
  Summary

## Slide 3: Introduction
  SCU and measurement system development timeline at KIT
  SCU 15 (superconducting undulator with 15 mm period length), served as light source for IMAGE Beamline (2015 - 2016)
  Test coils characterized in CASPER I cryostat at IBPT, long coils characterized at CERN
  SCU 20 (SCU with 20 mm period length), serves as light source for NANO Beamline (since 2018)
  Test coils characterized in CASPER I cryostat, long coils characterized in CASPER II measurement setup at IBPT
  Magnetic characterization of HEX70 wiggler for BNL (2021)
  Measurements performed in CASPER II
  Magnetic characterization of SCU16 for ANSTO (2021)
  Measurements performed in CASPER II
  Remeasuring of field integrals and QA onsite at BNET, Würzburg
  and check at ANSTO after delivery to Australia (2022)
  Measurements performed with mobile field integral measurement system
KIT and Bilfinger Nuclear & Energy Transition GmbH develop, manufacture, and test superconducting insertion devices (SCIDs) for the IBPT accelerators and low emittance light sources.

## Slide 4: Introduction
  Timeline…continued
  Magnetic characterization of SCU18 mock-up for S-PRESSO (EuXFEL, 2023)
  Tests performed in CASPER I and CASPER II at IBPT
  Magnetic characterization of THz-SCU for FLUTE at IBPT (2024)
  Tests performed in CASPER II at IBPT
  Magnetic characterization of SCW48 wiggler for DIAMOND (2025)
  Characterization at BNET (Würzburg) with local and integral
  mobile magnetic measurement system
  Magnetic characterization of SCW48 wiggler for ANSTO (2025)
  Characterization at BNET (Würzburg) with local and integral mobile magnetic measurement system
  Remeasuring of SCW48 field integrals onsite at ANSTO (2025)
  after delivery to Australia
  Measurements performed with mobile field integral measurement system

## Slide 5: CASPER I
Measurement setup to train and characterize the local magnetic field distribution of short superconducting coils (LTS and HTS)
  Applications
  Perform magnet training and quench tests
  Test new winding schemes
  New superconducting materials and wires
  New field correction techniques
  General
  Operating vertical, coils in LHe
  Maximum coil dimensions 35 cm length, 35 cm diameter
  Instrumentation
  Keithley constant current source (Hall sample current)
  Keithley multiplexer voltmeter (Hall voltage)
  Keithley Nanovoltmeters (HTS tests)
  2x 1500 A/+-5V power supplies to provide operating currents
  Quench detector
  Data logging system for quench analysis
KIT

## Slide 6: CASPER II
Measure magnetic field distributions and field integrals of superconducting coils with dimensions like in “real” IDs
(e. g. up to 2 m length, ~50 cm diameter, conduction cooled, arrangement horizontally)
  Magnet Training
  Quench detection
  Quench analysis
  Local field measurements
  3 Hall probes
  Field integral measurements
  Moving wire
A. Grau et al., IEEE Trans. on Appl. Supercond. 9001504 22-3 (2012)

## Slide 7: CASPER II
260 consecutive measurements show the reproducibility of the 1st vertical field integral measured with moving wire.
B
  Local magnetic field measurements
  Hall probes mounted on a sledge moving along the undulator length sliding on precisely machined guiding rails
  Hall probes calibrated to ± 90 µT (PPMS System at Institute for Technical Physics)
  Calibration for field -5 to +5 T
  Longitudinal position measured with a laser interferometer with sub-µm resolution
  IH Keithley constant current source, UH Keithley multiplexer multimeter
  Position measurement via laser interferometer
  Field integral measurements
  CuBe wire with a diameter of 125 µm is tensioned between two piezo stages placed at a distance of ~2.8 m (strain force 6 N)
  Coils magnetic plane to wire movements measured at room temperature. Alignment to 50 µm.
  Induced voltage measured by Keithley Nanovoltmeter after preamplification
  With FEMTO low noise voltage amplifier
  Reproducibility of I1 is ± 2x10-6 Tmm and of I2 is ± 1x10-5 Tmm2
  Local and integral field measurements can be performed within the same thermal cycle !!!

## Slide 8: Mobile Moving Wire Setup – Why ?
Improve stationary field integral measurement system to mobile version
  Insertion device is manufactured to operation.
  → No additional uncertainty occurs after cold tests of coils, coils shipment and implementation in the final cryostat (e. g. coil alignment)
  Cross check results from measurements in stationary cryostat possible (CASPER II)
  Mainly independent from final device length (depends on moving wire material and tensioning structure)
  Applicable to all vacuum chamber dimensions (small gap devices)
  Setup can be used to perform QA after device shipment (e. g. coils movement inside ID vacuum vessel during shipment) worldwide by comparing FAT results with SAT results.
  Because of local magnetic conditions compare results for multipoles and not absolute integral values
  Safe shipment of setup is essential to ensure measurement reliability
  To safe time for projects the setup has to be installed before cooldown and/or training
  In case of insertion device failure longer repair times occur because measurement setup has to be removed and installed after repair
  → This is less time consuming in stationary setups

## Slide 9
Mobile Moving Wire Setup - Layout
Moving wire
upstream
Moving wire
downstream
ANSTO SCU16 insertion device
FEMTO preamplifier
Keithley Nanovoltmeter
Controller turbopump US
Controller gauges US + DS
Controller turbopump DS
Controller strain sensor wire
Moving wire
SCU vacuum chamber
Moving wire tensioning
(Piezo actuator)
Piezo stage for horizontal movement
downstream, upstream
(vertical field integral)
Piezo stage for vertical movement
downstream, upstream
(horizontal field integral)
Signal connection
SCU vacuum chamber
Moving wire
upstream
Moving wire
downstream
Moving wire strain gauge
Moving
wire
Signal connection

## Slide 10: Mobile Moving Wire Setup
Components and Instrumentation
  2x Vacuum chamber Ø 250 with 4 access ports, connected to electron beam chamber (liner) CF 100
  Turbopump and prepumps upstream and downstream (liner pumped via setup chambers)
  Pressure reached in setup chambers ~7*10-7 mbar
  Bellows and fixing structure to ID to avoid alignment change during pumping and cooldown
  Measurement modules aligned with laser tracker to coils middle plane (+/- 25 µm)
  Piezo stages travelling range horizontally 60 mm, vertically 15 mm
  Signal voltage preamplification by FEMTO low noise amplifier DLPVA
  Voltage measurement by Keithley Nanovoltmeter
  Wire length 3587 mm (for this tests)
  Wire tension adjustable by Piezo-Mike linear actuator (upstream)
  This was needed because the wire shrinks during cooldown an did ripped off
  Wire tension gauge downstream
  After cooldown rise wire tension to 10 N for measurements
  Applicable to all liner chamber dimensions (small gap devices)
  Setup can be used to perform QA after device shipment worldwide by comparing FAT results with SAT results
  Comparison of fitting result for multipole component (normal quadrupole) gained during FAT in Würzburg and SAT in Melbourne fit quite well.
  FAT at BNET, Würzburg
  SAT at ANSTO, Melbourne
  At SAT, upstream and downstream part of the measurement setup were inverted. Therefore, position values and fitting slope is inverted.
Results of first application
  Absolut integral values are not nicely comparable because of different magnetic environment e. g. at FAT and SAT.

## Slide 11: SCID Mobile Magnetic Characterization Setup
Improve mobile field integral measurement setup with local field measurement components
  Advantages
  Insertion device is manufactured to operation.
  → No additional uncertainty occurs after cold tests of coils, coils shipment and implementation in the final cryostat (e. g. coil alignment)
  Field integral and local field measurements mainly independent from final device length
  Measurement lengths not restricted to ~1950 mm like in CASPER II
  Applicable to vacuum chamber heights of min. 5 mm
  (depends on HP sledge height and liner cross section layout)
  Setup can be used to perform magnetic characterization directly after manufacturing. No measurement campaign on coils in stationary setups
  Setup can be used after device shipment by comparing FAT results with SAT results.
  Setup has to be installed before cooldown and/or training
  In case of insertion device failure longer repair times occur because measurement setup has to be removed and installed after repair
  Challenges…How to
  combine both measurement types in one setup…inspired by CASPER II?
  design HP sledge for smaller vacuum gaps (height, length)?
  shape the sledge to incorporate all needed components?
  enable field measurement with Hall probe for roll-off evaluation?
  ensure proper sledge movement (pulling wire) and position measurement for long measurement ranges (> 2 m)?
  deal with the feedthrough of all signals. Moving wire, HP signal while sledge movement?
  move sledge into/out from liner for field or field integral measurements?
  park the sledge for field integral measurement (pulling wire)?
  park moving wire during HP measurements?
  ensure sledge moving limit in order not to harm any other component?
  keep all components at least passably accessible during installation?

## Slide 12: SCID Magnetic Characterization Setup
Chamber
downstream
Chamber
upstream
FEMTO
preamplifier
Keithley
Nanovoltmeter
Controller gauges
US + DS
Controller
Turbopumps
US + DS
Controller strain
sensor wire
Keithley
Multiplexer
Voltmeter
Controller Moving wire
Test liner
Temperature
readout
Prepumps US + DS
Laserinterfero-meter
Motor sledge
movement US
Motor sledge
movement DS
Guiding rails sledge
Signal cable (tape)
field measurement
(on bobbin)
Piezo stages moving wire US
Piezo stages moving wire DS
Limit switch
for sledge
Single beam laser interferometer head
Control and signal feedthroughs
Control and signal feedthroughs
Pulling wire sledge

## Slide 13: SCID Mobile Characterization Setup
  Moving stretched wire
  CuBe wire, Ø 125 µm, tensioned between two x-y piezo stage assemblies
  Length customizable to device length (latest test, ~3.8 m)
  Strain force adjustable until breaking by tensioning module (piezo motor)
  Strain force typically ~10 N
  Coils magnetic plane to stretched wire position measured at RT. Alignment to ± 25 µm.
  Induced voltage measured by Keithley Nanovoltmeter
  Preamplification with FEMTO DLPVA low noise voltage amplifier.
Piezo stage horizontal movement
Piezo stage vertical movement
Pulling wire sledge
Wire tensioning module
Piezo actuator
Moving wire
Signal connection
Pulling wire, sledge
(parking position)
Piezo stage horizontal movement
Piezo stage vertical
movement
Moving wire
(measuring position)
Sledge guiding rails
Stick with pin to shift Hall sample
(roll-off)
Strain gauge
Signal connection

## Slide 14: SCID Mobile Characterization Setup
Hall sample
ASensor HT144
3mm x 5mm x 0.5mm
Pulling wires
(Kevlar)
  Hall sample ASensor HT144 (4 wires), calibrated at low temperatures
  Calibration in PPMS system to ± 90 µT
  Hall sample glued to additional small sledge to gain roll-off data
  Hall sample shifting with stick, pin and moving wire piezo stage
  Temperature monitoring on sledge (CERNOX, 4 wires)
  LED attached (2 wires) to find retroreflector for laser adjustment along ~3 m
  Tape signal cable (10 strips), length ~5 m
  Tape cable coiled on spring loaded bobbin (sledge movement)
Tape signal cable
(10 strips)
Bobbin for tape cable
Connections for HP signal wires, T-sensor, LED
Wheels, bearings sledge movement
LED
Small sledge with HP
(shiftable)
Wheels, bearings sledge movement
Retroreflector
Connection
to tape
Interferometer pointing on retroreflector
Hall sample shifting
  Hall probe sledge
  Maximum sledge height 4 mm (suitable for vacuum gaps of 5 mm)
  Sledge width ~60 mm, to fit between vacuum chamber height
  Movement guided by precision rails, wheels and bearings
Sledge height 4 mm
Liner height 6.5 mm
Wheels,
guiding rails

## Slide 15: SCID Mobile Characterization Setup
Motor sledge movement
Deflection pulley for pulling wire
Laser interferometer on adjustment table
Pulling wire
Motor sledge movement
Pulling wire
Tape cable
Bobbin
Signal connections

## Slide 16: SCID Mobile Characterization Setup
  Alignment of vacuum chamber US and DS with laser tracker to insertion device fiducial points coordinates (± ~50 µm)
  Final alignment of guiding rails from vacuum chamber into liner by pins fitting to Ø 2.5 mm precision holes in front surface of liner support rails
  Support rails serve as guiding rails for the sledge while moving in the vacuum gap during field measurements
  Finally seal the vacuum chambers to the liner by closing CF 100 flange
  Insert pulling wire and clamp it to the sledge
  Insert moving wire and connect it to the piezo stages
  Connect moving wire signal
  Adjust limit switch to measurement distance
Moving wire in parking position

## Slide 17: SCID Mobile Characterization Setup – It works

## Slide 18: SCID Mobile Characterization Setup
First application, FAT (BNET) SCW48 for DLS, period length 48 mm, vacuum gap 8 mm, measurement length 1600 mm (4.3 T, 520 A)
  Because no Helmholtz coils were attached, it was not foreseen to correct the first (I1y) and second vertical (I2y) field integral, Auxiliary coils were not used for correction
  No correction for absolute field integral values at beam position
  I1y and I2y have to be corrected after commissioning in the ring, I1x and I2x can be corrected by outer horizontal steerers
  Main interest in multipole components
  Peak magnetic field should be reached
  Peak magnetic field not vary more than 1 % for each current
  Measurement tasks fulfilled

## Slide 19: SCID Mobile Characterization Setup
15/05/26
19
FAT (BNET) SCW48 for ANSTO Period length 48 mm, vacuum gap 6.5 mm, measurement length 2200 mm (4.5 T, 500 A)
  Similar tasks as for DLS Wiggler (No corrections applied)
  Found large horizontal field integrals
  BNET and ANSTO decided to deal with this during SAT (e. g. mount horizontal corrector along the whole device and remeasure
  Measurement tasks FAT, fulfilled !!
  Peak magnetic field should be reached
  Peak magnetic field not vary more than 2.5 % for each current
  Evaluate multipole components to be remeasured after shipment to Australia

## Slide 20: SAT at ANSTO with Moving Wire Setup
15/05/26
20
  Remeasuring of field integrals worked well
  Evaluation of multipoles show similar values as at FAT
  Calculations by BNET/ANSTO showed, that the large horizontal field integral FAT/SAT can be corrected for each current by mounting two outer windings along the wiggler
  The windings should reduce the horizontal field integral component by a factor of ~3
  Installation of additional windings by ANSTO personnel
  Reduction of horizontal field integrals I1y and I2y as predicted
  Measurement tasks SAT, fulfilled !!

## Slide 21: Summary
  Due to the long-term collaboration with BNET, the superconducting undulator projects at KIT brought significant impact to SCU technology in terms of device development
  Moreover, while and after the devices for the KIT synchrotron a lot of effort is put establishing stationary cryogenic measurement setups
  Leading to many measurement campaigns for KIT SCU devices and test coils (LTS, HTS) on the one hand
  and service measurements for and together with collaborating partners on the other hand
  In this regard KIT focused on the development of mobile measurement systems (1st setups worldwide) to gain the ability to characterize longer devices, everywhere
  Beside other applications the mobile moving wire measurement system is used to measure (FAT), remeasure (SAT) field integrals and evaluate multipoles to ensure the device operation after shipment
  Mobile setup for local field characterization and field integrals enables complete FAT/SAT magnetic characterization of final manufactured devices (location-independent) exceeding e. g. length restrictions of stationary setups (CASPER II)
  Results of both mobile measurement setups are fitting nicely
  Improvements/refinements of the setups are envisaged, and we are open for discussions/suggestions
  (implementation of pulsed wire, sledge design…)

## Slide 22: Thank you !!!
KIT