## Slide 1
QCM Accelerator Structure ESRF CDR Contribution
21 April 2026
Galen Aymar, Emilio Nanni, Andy Haase, Zenghai Li

## Slide 2: Agenda
QCM Overview
Schedule
Accelerator RF Design
QCM Design and Build Status
Alignment Requirements
External Mounts
Cryogenic Requirements
Vacuum Requirements

## Slide 3: QCM Overview
Original goal: demonstrator for C3
Additional goal: provide QCM to ESRF
  The full C3 design works for ESRF, but ESRF has less stringent requirements
VESSEL
ACCELERATOR STRUCTURE
QUADRUPOLE (not used)
RAFT
INSTALLATION HARDWARE
MOUNTS

## Slide 4: Proposed Schedule
Mechanical assembly of the QCM to start by Nov 2026, allowing testing before shipment to ESRF by July 2027

## Slide 5: Procurement Status
Long Lead Items
Accelerator Structure
Request for Proposal process completed via Sam.gov website – received 8 bids
A vendor has been identified, but paperwork and contracts are still in process
Lead time for the first set of parts is 28 weeks and for the complete two sets of parts is 34 weeks.
  Receive at SLAC in November 2026
Cryomodule
Received at SLAC in 2025

## Slide 6: Design and Build Status – Accelerator Structure

## Slide 7: Final RF Geometry
1-m structure: 40-cell with damping slots
Features
  40 cells with pair wise side to side distributed coupling
  Cell Balancing features
  4 damping slots extending to the beam center line (4 part open cells and iris)
  Structures to be spaced so that cell center to center distance is (N+.5)*26.242 at operating temperature
0.3mm
0.1mm
42.020 mm
21.010 mm
Damping slot
0.25 mm rounding
R2mm
Waveguide corner rounding

## Slide 8: Final RF Geometry
Position and RF Phasing Two or More C-Band Structures to Cancel Backward Acceleration
C-Band PI-mode cell length Lcell = 26.242 mm
Separation required to cancel acceleration: dL
  Measured from center to center of two end cells
  dL = (n + 1 + 0.5)*Lcell
  n = 0, 1, …
Structure RF phase
  phase_down – phase_up = -(n+0.5)*180 degree
upstream structure
(phase_up)
downstream structure
(phase_down)
Beam direction
dL
Note: all dimension measured at operating temperature
  applies to more than two structures

## Slide 9: FEA: Deflection Under Pressure, 2D
The “2r” wide damping slots create continuous vacuum space that must be supported to prevent excessive displacement of the cell surfaces. Small predictable elastic deformation can be compensated for during structure tuning.
This was initially reviewed with simplified geometry in 2D to determine if the span would need support within the damping slots or if the full span could be supported by the structure.
With distortion in the range of .3µm conclusion is that the damping slots can remain continuous without interruption for support as long as the inner WG joint is well connected.
The motion of the upper end of the cavity wall under external pressure in microns is:

## Slide 10: FEA: LN Pool Boiling
Chart shows heat transfer coefficient in nucleate and film boiling regimes as well as transition temperatures. The design goal is to avoid film boiling, staying in nucleate boiling.
To remain in nucleate boiling (at nitrogen pressures between 1 and 2 atm) the surface/bulk temperature difference must remain below ~10 K.
In nucleate boiling the heat transfer coefficient will vary from ~.1 to 10 W/cm^2 as the surface/bulk temperature difference ranges from ~1 to 10K. (.1 to 1 W/cm^2/K)
Other sources were checked with similar coefficients. Simulations indicate that this is not critical as the surface heat flux is well below onset of film boiling.

## Slide 11: FEA: Structure Cooling
Our typical (and proposed) tuning features create 2 potentially problematic conditions.
Gas trapped areas that will not be well cooled
Areas close to the cell walls that may flash to film boiling
Used 12 cells and included “end effect”
Loss then 2500/40*12/4=187 > round up to 200 W
Cooling ranged from all exterior to just the main faces
Used convection coeff from ~relevant sources

## Slide 12: FEA: Pool Boiling Temperatures
Case with no cooling at the tuner surfaces indicates inside temperatures at approximately 4° above the LN bulk temperature.
As this is “well below” the film transition temperature of ~10°, no modification of the tuning features is thought essential.
This also shows the extent of the additional end cooling which is discussed on subsequent slides.

## Slide 13: FEA: Heat Transfer, Pressure, & End Effects
Distortion from RF heating and 3 atm pressure (1.5x design)
Case without any tuner surface heat transfer
Displacement of less than 0.5 µm is acceptable
Some tuning offsets to compensate for end effects may be incorporated.
Deformation normal to the damping slot

## Slide 14: FEA: Heat Transfer, Pressure, & End Effects
Von-Mises stress below 1 kpsi. Annealed copper yields at >3 kpsi.
Yielding is not expected
Conditions from previous slide
Distortion 3D visualization

## Slide 15: Mechanical Features - External
Part handling feature for machining – TBD by vendor
Mounts for Rasnik or laser nests
Braze clamping pads
Cell tuners
Thru-holes for bolt clamping at braze
Mounts for struts to support raft
Alignment surfaces to be machined in one pass at the final step, checked/recorded

## Slide 16: Mechanical Features - Internal
Braze surfaces
Waveguide feed
Damping slots
Cells and irises
Internal radii around cells at split face
Riris
Rcell
Rdamping slot

## Slide 17: Braze Steps and Features
3 critical “perimeter” joints (2 vacuum, 1 mechanical)
During braze, the braze joints are held by external clamps and compression bolts.
At assembly, alignment is achieved with a pin and slot arrangement, using stainless steel inserts to prevent deformation of the hole and slot (exterior flat surfaces may also be used)

## Slide 18: Design and Build Status – Cold Box/Fill Lines
Cold boxes designed to enable controlled filling/venting of cryomodule.
P&ID agreed upon with SLAC pressure vessel safety team.
Valves are already received at SLAC.
Half of the transfer lines and bayonets are received at SLAC, but will require additional components or modification.
Boxes and lids are in final design.

## Slide 19: Design and Build Status – HPRF Waveguide
Feedthrough into the QCM is via stainless steel waveguides bolted externally and internally.
Internal waveguides are of copper and have an elongated path to make them flexible, allowing for mm-scale adjustments of the raft and accelerator while the feedthrough is fixed.

## Slide 20: Design and Build Status - Feedthroughs
Proposed feedthrough locations
Vacuum
Alignment/ unused
Alignment/ unused
Service Hatch
HPRF & N2 Vapor Out
LN2 In
Instrumentation
LN2 Drain
Mechanical Centering Pin
Beamline

## Slide 21: Design and Build Status – Mounts/Movers/Loaders
5 adjustable mounts for 5 degrees of freedom
Adjustable range is mm-scale, with the exact limits to be agreed upon with ESRF and then set by SLAC.
Animation showing various movements for one pair of movers
1
2
3
4
5

## Slide 22: Mechanical Features – Alignment and Fiducials
ESRF needs: Precision alignment of the structures in the QCM to the machine coordinates (beam tunnel/mating components) and a means of verification/adjustment.
Will want to check alignment while warm with cryovessel open – confirm if any movement during shipping.
Needs confidence of the location while cold, however that is established.
Option 1: While warm, transfer datums to external fiducials and use a calculated or empirical “cold-warm” adjustment.
  Need to qualify and ensure consistency in this adjustment factor.
Option 2: Use the Rasnik alignment system to be installed and tested by SLAC.
Option 3: Use optical viewports through vacuum and liquid nitrogen to directly measure fiducials on the accelerator.
  Unclear if this is even possible and requires additional experimentation.

## Slide 23: Mechanical Features – Alignment and Fiducials
ESRF: Finalize mechanical alignment accuracy required before switching to beam-based alignment.
SLAC: Perform small-scale tests of measurement through windows and LN2.
SLAC: Investigate alternative alignment methods.
SLAC: Propose a plan for qualifying the alignment method chosen to ESRF’s needs.
Actions

## Slide 24: External Mounts
Adjustable mounts were originally designed by Fermilab for the LCLS-II LINAC cryomodules.
Allow for vertical, transverse, and longitudinal adjustment.
Include additional seismic restraints on the sides.
Vertical
Transverse
Longitudinal

## Slide 25: Cryogenic and Vacuum Requirements
LN used on cooling fill: 1400 L
  Vessel volume: 550 L
  Cooling the structure: ~720 L
  Loss during fill: ~100 L (over 12 hours)
Standby losses: 133 L/24 hr for 286 W standby power loss
  Radiation: ~247 W
  Conduction: ~38 W
  Convection: 0 W (assuming <1x10-5 torr vacuum)
Operating losses will vary and be in addition to the above.
Vacuum
  Primary concern is to maintain vacuum <5x10-5 torr to ensure no convection heat loss.
  This vacuum level requires turbo pumps (and appropriate backing pumps), but the exact size of pumps depends on system outgassing and leak rates, which will be determined during testing at SLAC.

## Slide 26: Conclusion