## Slide 1
COLD Project Review Meeting
WP2 – Diagnostics & Laser for Photo Gun
Friederike Ewald
On behalf of the EBS accelerator team and the collaboration: SLAC, INFN, CERN, PSI, Cornell

## Slide 2: Outline
T. Brochard/MEG
Preliminary 3D-sketch of the COLD setup
Part I – Laser and optics for the photo gun
  Requirements
  Technical choice
  Implementation
Part II – Diagnostics
  Diagnostics layout for phase 0
  Diagnostics layout for phases 1 & 2
  Specifications

## Slide 3: Part I – Drive laser for the RF Photo Gun
Requirements / Specifications
Possible candidate for the laser
Synchronisation to RF
Layout of laser installation
Conclusion

## Slide 4: Requirements For RF-Photon-Gun Drive Laser
Photo cathode material | Copper
Quantum efficiency QE for electron emission from Cu: | ~10-5 – 10-4
 Work energy for Cu-cathode | 4.53 – 5.1 eV
 Laser wavelength in the UV | < 243 nm – 274 nm, e.g. 266 nm from 3rd harmonic of Ti:Sa e.g. 257 nm from 4th harmonics of Yb-based lasers
Electron bunch charge (wished) | ~ 1 nC
 needed Laser pulse energy (4.3 eV * 1 nC/QE) | ~ 50 – 100 mJ on the cathode Losses during beam transport:  200 mJ minimum at laser output
Parameters defined by RF-cavity and LINAC design | 
Laser pulse duration | ~200 fs
Laser spot shape and size on the cathode | 600 mm radius, round flat top (maybe Gaussian OK)
Longitudinal pulse profile | Gaussian
Other requirements | 
Repetition rate: | 4 Hz… 10 Hz
Synchronisation capability | Phase locking to RF (2.9989 GHz & 5.712 GHz)
Stability | High reliability and as much as possible maintenance free
Limited Space | Compact dimensions

## Slide 5:  Yb solid state laser technology
Good candidate: Pharos 2H-4H from Light Conversion
Diode pumped Yb-crystal laser (l = 1030 nm) with integrated 4th harmonics generation:
  l = 257 nm
  E = 200 mJ / pulse (*)
  190 fs pulse duration minimum
  Transverse beam diameter at exit ~7mm
  Single shot – 100kHz
  Oscillator frequency 60 – 84 MHz
  Compact
This laser is already used for RF photo guns at
DESY (S-band photoinjector at SINBAD-ARES)
CERN CLEAR/CTF2 photoinjector
Photoinjector of CXFEL (Arizona State University)
827 cm
492 cm
25 cm
Positive feedbacks on performance and maintenance (beam quality, stability, long term reliability,…)
190 fs
(*) 2 mJ at fundamental wavelength (1030 nm) 10% conversion efficiency into 4th harmonics (257 nm)

## Slide 6: Synchronisation of Laser to RF
The laser is driven by a mode locked oscillator.
The oscillator frequency may be chosen within certain limits (for the Pharos: 60 – 84 MHz)
The common RF-frequency between European S-Band and C-Band:
142.8MHz
Gun
2998.8 MHz
Accelerating structures
5712.0 MHz
x 21
x 40
Laser oscillator frequency
71.4 MHz
÷ 2

## Slide 7: Approximate Schematic Principle of the Laser Installation Layout
Photo gun
LINAC tunnel
Modulator hall (Pinj)
Laser cabin (temperature stabilized ± 1 deg C)
Laser
Electrons
Chiller
Optical table
with enclosure
Optical breadboard
LINAC tunnel wall
Shutter
Contact switch
UV-viewer for beam spot observation
To be defined
To be defined

## Slide 8: Laser Beam Transport
Wish list:
Distance from laser output to cathode as short as possible
Adjustment of beam dimensions (horizontal / vertical) on the cathode surface (~300 um radius)
Adjustment of transverse position on the cathode
Adjustment of photon flux on the cathode
Maybe adjustment of pulse duration
UV wavelength and short pulses :
Damage thresholds of optical coatings and bulk material to be taken into account
Fused silica for transmission optics and viewports
Carbon deposit on optical surfaces will affect photon flux and beam profile quality
Total beam losses due to beam shaping and transport can be very high
Choose laser pulse energy enough margin
If pulse stretching is needed, this will come at the cost of pulse energy
Where possible use (dielectric) mirrors instead of lenses
As few optical elements as strictly necessary
Is the flat top transverse profile really needed? (ARES & CTF2 experience: No)

## Slide 9: Part I – Conclusion on Drive Laser for Photo Gun
Specifications of the Pharos 2H-4H laser meet perfectly our requirements
Positive feedback concerning operation and reliability from other institutes
Why not Ti:Sa technology?
  As compared to FEL LINAC installations we do not need complicated time-structures, which would impose Ti:Sa technology
  Much higher space requirements, lower reliability/stability during operation, more maintenance, more expensive
Next steps are the design of the optical pathway / beam transport up to the cathode surface
Thanks to F. Burkhart & W. Kuropka (DESY / ARES) and E. Granados (CERN / CTF2) for discussion and advice

## Slide 10: Part II – Electron Beam Diagnostics
Approximate layout of diagnostics installation for phase 0 of the project
Approximate layout of diagnostics installation for phases 1 & 2 of the project
Diagnostics specifications
Conclusion

## Slide 11: PHASE 0 – Accelerating Structures Without Photo Gun  Dark Current Measurements
Beam Dump Thom-X (50 MeV), Soleil 150 MeV design
Electron beam
Lead
Graphite
Neutron absorber
Expected dark charges in the order of nC
Faraday cups is the beam dump.
Two Beam Dumps for 350 MeV (maximum)  to be designed by Radio Protection Group
A screen may complement the diagnostics in Phase 0.

## Slide 12: PHASES 1 & 2
Phase 1: Photo gun and the diagnostics & dump section will be installed, leaving space for the installation of a shorter accelerating structure in Phase 2.
The standard diagnostics for phase 1 and 2 are essentially the same.
Phase 2:
RF-deflector for longitudinal profile; depending on budget

## Slide 13: PHASE 1 – Installation and Commissioning of Photo Gun  RF-Gun Diagnostics
Photo Gun
Solenoid
Steerer
Current transformer
Laser
BPM
Screen monitor
3D-view of Gun Section (from INFN):

## Slide 14: PHASE 1 – Installation and commissioning of photo gun  Diagnostics Section
Insertable Screen monitor for
Beam position and
Transverse size (1 – 3 mm)
YAG:Ce or GAGG:Ce with adequate imaging for high resolution
In-flange ICT (commercial solution)
Beam charges ≤ 700 pC
 Stripline? Button BPM? Cavity BPM?
 Which resolution is needed / achievable? 2 ps, single-pass 4 Hz, beam charge ≤ 700 pC
Dipole magnet + Screen + Imaging optics
Specified for maximum of 350 MeV
(for Safety, beam max. energy = 300 MeV)

## Slide 15: Diagnostics Specifications Summary
Beam parameter | Order of magnitude | Instrumentation | Remarks
Electron energy Energy spread | 0 – 350 MeV 0.18 % | Dipole spectrometer with scintillating screen + imaging | Dipole design  IDM Group with specifications from Beam Dyn.
Dark current during commissioning | ≈ nC | Faraday cup (FC); Integrated into beam dump | Beam Dump design  Safety / RP group
Bunch charge | ≤ 700 pC | FC; Current transformer (ICT) | Commercial solutions available
Transverse beam profile | ≈ 1 – 3 mm | Screen monitors (destructive) Optical Transition Radiation (OTR) ? | Screen monitor resolutions < 5 um possible
Emittance | 5.5 mm.mrad | Focus scan + screen monitor | 
Beam position | ≈ mm … ? To be specified | Stripline/ Button-BPM / Cavity BPM / Screen  Choice of BPM-type still open | Achievable single pass resolution(RMS) in the order of 10 mm
Bunch length | ≈ 2 ps | Deflecting cavity | 
https://doi.org/10.1063/1.2401409
https://doi.org/10.1016/j.nima.2006.07.050

## Slide 16: Part II – Electron Beam Diagnostics - Conclusion
Required beam parameters and general layout still evolving
Very early stage of diagnostics planning, few details yet
We will mainly rely on
  Well established techniques / instruments already used at the ESRF
  On the experience / instrumentation of other LINAC installations
Visit at CTF2 & CLEAR at CERN in November 2025
Visit at PSI January 2026
Contact with ARES/DESY in 2025
Collaboration INFN Frascati Visit at MAXIV …
Screen monitors
CTF2/CERN
CLEAR cavity BPM
CTF2 Spectrometer screen
PSI RF-deflector
Dipole CTF2/CERN

## Slide 17: Human resources for WP2
 | Beam Dynamics Group | Diagnostics Group | Mechanical Engineering/ Drafting (ESRF) | Building & Infrastructure (ESRF) | IDM Group (Magnet design) | Other
Optical layout for drive laser installation |  | X | optics engineer |  |  | 
Laser room Hole drilling |  | X | X | X |  | 
Laser installation/ maintenance |  | X +? | ? |  |  | ?
BPMs | X | X | X | X |  | Depending on type of BPM  fixed term contract?
Dipole spectrometer | X | X | X |  | X | 
Screen monitors | X | X | X |  |  | 
Current transformer | X | X | X |  |  | Commercial product
Faraday cups |  | X | X |  |  | Radioprotection group For beam dump design
Work load on the long term?

## Slide 18: Closing Remarks
Laser requirements well defined.
Rough needs and layout for beam diagnostics instruments are defined
Precise specifications, technological choices, simulations and design of diagnostics to be done
Principle of choice of equipment (diagnostics as well as laser installation):
  As simple as possible, without compromising the needed performances
  Reliability and low maintenance needs
  Based on:
      commercial instruments,
      in-house knowledge and
      experience from other institutes (BPMs, screens monitors,…)
Most of the work can be done with in-house resources from ASD, ISDD and TID. We may need to consider additional resources for
  The long-term laser maintenance (in-house?)
  Maybe for the design of BPMs (if not standard solution) & during installation & commissioning phase

## Slide 19
Thank you for your attention.