## Slide 1
ASD-day, Grenoble, January 24, 2023
Kees Scheidt
beam stability issues and
diagnostics developments :

## Slide 2
ASD-day, Grenoble
January 24, 2023
on behalf of the
Diagnostics Group :
Nicolas Benoist
Elena Buratin
Guillaume Denat
Friederike Ewald
Benoit Roche
Kees Scheidt
Fouhed Taoutaou
Kees Scheidt
beam stability issues and
diagnostics developments :
the Diagnostics Group

## Slide 3
the beam’s positional stability
too much reliance on the 320 e-BPMs only  can NOT warrant a perfect stability ...
version-1 X-BPMs on two BM-sources : full results >1year, but (almost) only during MDT
version-2 X-BPMs on two BM-sources : now (Jan. 2023) installed  first results & perspectives
temperature stabilization in the cubicles of the BPM-electronics  in good progress
the upgrades of the (slow/fast) orbit control system
Diagnostics Developments
two units of non-destructive vertical beam Halo monitors  excellent results
the measurement of ultra-low beam-current, down to the single electron
general upgrades, activities, corrections or improvements on numerous Diag. systems
1
2
3
4
6
7
8
5
outline of this presentation

## Slide 4
Beam’s positional stability : How good is it really ??
“the EBS has very good stability ”
“the beam stability is excellent ”
“the stability is really perfect !! ”
image from the
X-ray pinhole camera
it is an image, NOT a movie …
if I had put a movie
you would NOT see the difference :
the beam appears extremely stable !
this since the first days
of commissioning of EBS
i.e. 3 years ago
January 2020

## Slide 5
AC domain stability is very good  thanks to girders etc. etc.
Stability in the low-AC domain (1 – 100Hz)
old ring 2010, FOC On & Off new ring, no FOC (yet)
Hoff
Hon
H
V
Von
Voff
comparison between the old (data from 2010) and the new EBS ring :
freq. [Hz]
freq. [Hz]
from ASD-day 3 years ago, 2020

## Slide 6
stability needs assessment over a large time domain
time span
frequency [Hz]
1
0.1
0.01
0.001
10
100
1000
day
month
week
[sec]
10
10
4
5
hour
0.1
1
10
100
1000
AC domain (vibrations)
“slow” (movements, drifts)
previous slide
with our present e-BPM system we are :
excellent !!
un-certain
more difficult

## Slide 7
we are really good in the fast domains (faster then 1hr)
time span
frequency [Hz]
1
0.1
0.01
0.001
10
100
1000
day
month
week
[sec]
10
10
4
5
hour
0.1
1
10
100
1000
AC domain (vibrations)
“slow” (movements, drifts)
previous slide
with our present e-BPM system we are :
excellent !!
un-certain
more difficult
we see our beam
jumping every minute …
we have a 7Hz
in our beam …

## Slide 8
we need to better address (only) the slow domains
time span
frequency [Hz]
1
0.1
0.01
0.001
10
100
1000
day
month
week
[sec]
10
10
4
5
hour
0.1
1
10
100
1000
AC domain (vibrations)
“slow” (movements, drifts)
previous slide
with our present e-BPM system we are :
excellent !!
un-certain
more difficult
there is 1hr periodic
motion in our beam …
our beam is not
where it was
5 days ago …

## Slide 9
layout of the e-BPMs and the Steerers
L=fast-BPM (6) S=slow-BPM (4)
BM-
source
fast/slow
corrector
fast/slow
corrector
fast/slow
corrector
slow
correctors
slow
correctors
some 7th BPMs in the Ring
suffer from instabilities
ID-source
ID-source
the EBS orbit stabilization has distinctive FAST & SLOW parts
for the FAST part it uses 192 e-BPMs (later all 320) and 96 fast steerers , both H & V planes
for the SLOW part it uses 320 e-BPMs and 288 slow steerers, also both H & V planes
 in EBS we had NOT foreseen X-BPMs  so our e-BPMs are both Judge & Party

## Slide 10
X-BPMs on BM-8 and BM-16 added in Oct. 2021
L=fast-BPM (6)
S=slow-BPM (4)
BM-source
some 7th BPMs in the Ring
suffer from instabilities
6
7
SBM of 0.875T
the new
X-BPM
X-ray beam

## Slide 11
available Front-End chambers at 23m from source
X-BPM @23m
from source
BM-beam
radio-protection shutter
end of
Tunnel
out of 17 BM sources
11 Front-Ends have such chamber

## Slide 12
the V1-X-BPM for BM Oct-2021 to Dec-2022
X-ray beam
Scintilator
15 x 15 x 1.5
mirror
bellows &
motorized translation
optics & camera
simple and
low-cost device
water-cooled absorber
VERSION_1

## Slide 13
the V1-X-BPM for BM Oct-2021 to Dec-2022
in yellow :
the existing
FE-chamber
an image of the X-ray beam is
projected on the scintillator (green)
the light emitted from that is captured/focused
on to a camera via mirror (blue), window, lenses
a water-cooled absorber is in front of the scintillator
IN
OUT
camera
optics
UHV-window
so when the BM uses the beam then
this X-BPM is OUT (>15mm)
>15

## Slide 14
this X-BPM yields images we use only vert. position
image
Vert. profile
Vert. Position
of beam center
is interceptive  we can only use 1 day/week (MDT)
only yields vertical position  is not such on issue
is slow (few Hz data rate)  is not an issue
only vertical position
is of use & interest
this V1 X-BPM :

## Slide 15
comparing between two independent position results
L=fast-BPM (6)
S=slow-BPM (4)
BM-
source
6
7
SBM of 0.875T
Pos V = 0.8060 *V_BPM_6 + 0.0990 *V_BPM_7 Ang V = -0.9696 *V_BPM_6 + 0.49478 *V_BPM_7
the new
X-BPM
X-ray beam
comparing the vertical position of the new X-BPM
with that from the e-BPMs in the Ring
@23m

## Slide 16
discrepancy between two independent position results
BM-16
observation :
the X-BPM and e-BPM results drift in opposite direction, attaining 150um difference …
67 continuous hours in 20mA, 10-12 December 2020
in contradiction to what is so often said
about EBS positional stability ( “very good”, “excellent”, “perfect” )
we now have clear indication/proof that it is NOT …
1) what is the cause ? 2) what can we do about it ?

## Slide 17
the cause of the discrepancy : temperature of e-BPMs
strong correlation between BM-16 X-BPM and temperature of that C16-6 Spark
this measurement was only possible since BM-16 allowed us
to keep the X-BPM inserted for a full 6 days
Spark
temperature
X-BPM
Vert. position

## Slide 18
from Version-1 to Version-2  permanent monitoring
V1 :
the interceptive
X-BPM
served in 2022
V2 :
the fixed & NON-interceptive X-BPM
results since January
DQ1D
SBM
retractive
fixed !
DQ2C
courtesy of J.Chavanne
DQ1D
SBM
DQ2C
SBM hor. width inside X-BPM-chamber,
i.e. 51mm at 23m
fixed !
or
the un-used X-rays of close-by EBS dipoles
can be used with the same purpose

## Slide 19
replacing the V1 with the V2 X-BPM for BMs
VERSION_1
installed this
January on
both BM-8
and BM-16
VERSION_2
simpler, cheaper
more efficient !
see results next slides !

## Slide 20
What about Insertion Device X-BPMs ?
I wanna be
an X-BPM !!
the Front-Ends of many (11) IDs
also have an identical chamber
available at the very end of the FE

## Slide 21
the X-BPM for IDs would be interceptive & retractable
in yellow :
the existing
FE-chamber
IN
OUT
camera
optics
UHV-window
so when the ID uses the beam then
this X-BPM is OUT (>15mm)
>15

## Slide 22
What use for an Insertion Device X-BPMs ?
IN
camera
optics
UHV-window
1) very regularly different beamlines ask for (and get)
an angular bump (on the e-beam) so to bring their
X-ray beam back to where they want it
we have no way to know if such X-ray beam-shift was
caused by the EBS e-beam or by the beamline itself
with such an X-BPM we would know
2) the electronics of e-BPMs have excellent reliability
but can have a rare failure. Replacing them during USM
is sometimes unavoidable.
this can be a problem for those e-BPMs that are
around the IDs : this replacement can introduce a small
but very disturbing angular shift of the e-beam,
and thus ID X-ray beam.
the X-BPM would be used to compare new position
values with those taken every week (during MDT) and
can thus serve to correct such shift.

## Slide 23
situation of temperature fluctuations on e-BPMs (Sparks)
2 deg. between
winter/summer
1 deg.
fluctuation
day / night
1 deg.
fluctuation
day / night
these are averages of all 128 Spark units, the worst ones are a few times worse !
JUNE
FEBRUARY
2 days

## Slide 24
home-made
solution
Costs : 500 Euros
per unit (cubicle)
solution to dampen the temperature fluctuations of BPM
in total 17 units to be done,
5 units partly done in 2022
all 17 will be completely done by summer 2023

## Slide 25
installation in good progress, will be finished summer 2023 (17 units)
ACU has delivered an efficient device-server for the temp. control
C16  temperature control installed & active
5 other Cells
no control yet
5 days
Spark
Temperature
[deg. C]
C16

## Slide 26
first results with permanent V2 X-BPM on BM-16
blue=X-BPM-vert. position
red=temperature
0.2deg
46um
46um at 23m = 2urad !
Jean-Louis Hazemann of
BM-16 asks for 1urad stability …
the temperature control in SR-BPM cell 16
was ON but not well configured
due to suddenly cold weather of last week
2 days

## Slide 27
limitations of temperature control on e-BPMs (Sparks)
2 days
[C]
3 cells NO temp. control versus 2 cells WITH control
Cell-16
Cell-28

## Slide 28
situation of temperature fluctuations on e-BPMs (Sparks)
blue=X-BPM-vert. position
23um
=1urad
red=temperature
0.2deg
+/- 0.1 deg temperature control is the best we can do ! if not good enough
then we could use the X-BPM to further stabilize the BMs beam position
1 day

## Slide 29
conclusion and prospects for X-BPMs
out of 17 BM sources 11 Front-Ends have such available chamber :
7 x SBMs (Single-Bending-Magnet)
4 x 2PWs (2-Pole-Wiggler)
the costs per unit is about 8 KEuros
this year (2023) we expect to install about 4 SBMs
and the next year(s) the 7 others
out of all 17 ID sources 11 Front-Ends have such available chamber :
the costs per unit is about 12 KEuros
this year (2023) we intend to install one prototype (May, or August)
perhaps one more in 2nd part of 2023
and if prototype experience is successful then
the others in the next year(s)
BM
ID

## Slide 30
the upgrades on the fast orbit correction system
2) Optimization of the PID parameters
for more flexibility & performance
the old system for setting/controlling
these parameters has been fully and
drastically modernized
this allows a strongly increased efficiency
during MDT studies in assessing and
optimizing the performance and behavior
of the FOFB
application window
1) Adding the 10KHz data-stream on all 128 Spark-BPM-electronics for optimized
Fast Orbit Feedback (FOFB) and specific & dedicated applications

## Slide 31
two units of vertical non-destructive HALO-monitors
CH12
SR/FE
valve
CH11
CH14
vertical
movable
absorber
In-air Halo
detector
CH13
Alu window
X-rays
Components:
dipole source (DQ1D, 0.57T)
free beamport
vertically adjustable X-ray Absorber to shadow the intense beam-core
Aluminium window to let the X-rays from UHV to air
sensitive Detector to image the X-rays
in the (free) Front-Ends of cells 10 & 11

## Slide 32
two units of vertical non-destructive HALO-monitors
Hor.
Vert.
typ. 2mm
beam-core

## Slide 33
our Diagnostics now measure all kind of electrons
the dying
electrons
Halo
electrons
50nA
the “work-horse”
electrons
all our diagnostics
measure these
measured by
Lifetime and
beam-loss
detectors
no diagnostic before,
now two HALO-monitors
2mm

## Slide 34
the HALO-monitors are an excellent watch-dog on Vacuum
BeamLoss
detectors
the 2 Halo
monitors
vacuum
gauge

## Slide 35
single electron(s) in the EBS
at 200mA there are : 3 500 000 000 000 electrons
can we put & measure only 1 electron ?
offset
1st electron
3rd electron
2nd electron
4th electron
5th
offset
beam kill
typical time of 10min
is 2400 injections
integration at
8sec sampling
to reduce noise
time [minutes] , sampling at 8sec period
Detector Level [A.U.]

## Slide 36
other developments & activities
improvements on the damping of injection perturbation
this is a FeedForward system that often needs fine-tuning & optimization
the precise calculation of the FF parameters is now semi-automized
 see presentation Simon White
2) all emittance-monitors now have Injection-synchronized gating added
this allows to much better assess the effect of injection perturbations
and also the effect of its damping system

## Slide 37
ASD-day, Grenoble
January 24, 2023
The Diagnostics Group :
Nicolas Benoist
Elena Buratin
Guillaume Denat
Friederike Ewald
Benoit Roche
Kees Scheidt
Fouhed Taoutaou
Kees Scheidt
we have lots of essential systems to look after
and projects that will yield benefits to the ASD
we just need to keep our staff !!
KEEP OUR DIAG. STAFF !!

## Slide 38
a big thank-you for help & support
from colleagues in :
Vacuum , RF , Beam-Dynamics , ACU ,
FE , Operation , ALGE , TID
Radio-Protect. , Mech. Engineering
ASD-day, Grenoble
January 24, 2023
Kees Scheidt
on behalf of the
Diagnostics Group :