ASNR/Safety linac&Cold meeting:
10/07/2026
Present: Patrick, Paul, Jean-Luc (minutes), Simone, Marc, Thierry, Friederike, Alessandro, Yann
Excused: Qing, Simon
· Status ASNR for modulators (Klystons)
· 30 April update of the ESRF authorisation, 2 klystron S band klystron included and done.
· C Band in instruction
· Status ASNR for linac refurbishment 
· RF conditioning no need of authorization (only for klystron)
· Use of cavity spare to replace section 1 and 2 need authorization
· Status of document and request:  In instruction
· Expected date for authorisation  September/October 2026
· Status of ASNR for COLD
· Simone should send the latest version of beam profile
· Beam dump design is the critical issue
· In order to prepare the instruction Paul/Patrick could already send the CDR to ASNR for information (Simone dispatched the pdf just after the meeting)
· Estimation of the operational use of COLD. A preliminary planning is attached. For ASNR a factor 2 of margin should be at least included/

Paul needs:
· -updated stp file of the cryomodule (Thierry)
· stp file of the implementation of the COLD linac inside the tunnel (Thierry)
·  geometry + beam optics of the part between end of linac and beamdump
· - maximum accelerated electron charge per year for ASNR authorisation (ASD)
· - info about ground water level (Yann, sent it today).

· 350 MeV Beam dump design
· The beam dump is a critical issue for the study
· Paul presented a preliminary study of a beam dump (size, material, position) [attached file]
· The activation of the groundwater should also be studied even if it should not be critical
· Paul prefers for safety reason that the beam dump is located in the floor (ground)
· Dark current is as important as the main current for the design of the beam dump
· Dark current could be large at the beginning during conditioning
· Dark current could be a different energy and could reach nC
· Dark current should be simulated but it is difficult to do
· Beam dump is needed also in phase 0 for the dark current
· SLAC could help us.
· China did the same exercise at Heps. Qing could contact them.
· LN2 protection, pressurized vessel
· Type of leak to consider, worse case leak
· Volume to be consider for the blow away still to be defined
· Protection, extraction to be reconsider
· Patrick is preparing a document
· Meeting second week September with BIG/ Patrick/ Simone 
· CEA / SLAC could be contacted for advice and data if needed
· Infrastructure
· Holes in the floor for grounding, this summer shutdown
· Holes will be done in the linac tunnel during October and Winter shutdown
· Resource
· A PHD student will join the cold project in August

· Next meeting
· First week of September
